## Slide 1: HPD system development
April 22 2006
OUTLINE
HPD system requirements
HPD readout system R&D status
FJ collaboration R&D items
Details
Frontend electronics
Waveform sampler (Analog memory)
Digital filter
Plan

## Slide 2: HPD readout system requirements
HPD requires readout electronics on detector. (signal of HPD ~105 or less)
Readout electronics specification
  ～＜1nsec timing resolution
  Good S/N ratio (isolation of 1p.e. 2p.e. .. peaks)
Followings should be considered for large volume detector system
  Stable system (EMI, GND loop, etc)
    Ground isolation (optical device for data transmission etc.)
    Noise filtering system which has adaptation to environment
    …….
  R&D for timing distribution and control system for whole detectors distributed in wide area.

## Slide 3: HPD readout system development
System : Digital filtering on the HPD
Status :
  Details are described in page 5~14
    Frontend
      ASIC development almost finished.
    Waveform recorder
      prototype ASIC (record length=64nsec) is evaluated
      2nd ASIC (record length=512nsec) is under development
    Digital filter
      function confirmed using FPGA
    Data transfer and Control
      Discussion
    HV system on HPD

## Slide 4: France Japan collaboration items
Data acquisition system for large volume detectors
  Collaboration form： Discussion and exchange information each other for specifications and future development
  R&D for timing distribution and trigger system
  Discussion for readout system of HPD
    Waveform sampling scheme
    QVC & TDC scheme
  R&D for data transmission scheme from a detector to backend
Development of readout components
  Collaboration form：Discussion in France/Japan, Development in Japan
  Development of waveform sampler
  Development of 20kV high voltage power supply
System test（readout system with HPD）
  Collaboration form：System test in Japan.

## Slide 5: System Overview
Pre-amplifier (ASIC)
Digital Filter
(FPGA)
Waveform sampler (AMC)
Detector
(13-inch HAPD)
(Q, T)
A/D conversion
f
t
～＜1nsec timing resolution
1p.e.が分離できるアナログ性能を同時に満たす。
かつ環境雑音などにも強いフレキシブルなシステムをめざす。

## Slide 6: Frontend electronics
Pre-amplifier TEG (ASIC)
Digital Filter
(FPGA)
Waveform sampler (AMC)
Detector
(13-inch HAPD)
Low noise
Wide bandwidth
Wide dynamic range
f
t

## Slide 7: Preamplifier ASIC
Preamplifier TEG
Noise ~3000 electrons @Cd =100pF
Output : 2V max
Prototype ASIC Preamplifier
Impulse response（rise time <1nsec)

## Slide 8: Waveform sampler
Pre-amplifier (ASIC)
Digital Filter
(FPGA)
Waveform sampler (AMC)
Detector
(13-inch HAPD)
Fast sampling (~1GHz)
f
Low power
Wide dynamic range

## Slide 9: Analog Memory Cells (AMC) TEG
No fast sampling clock  low cost, low power
A trigger propagates the delay chain, turning off analog switches one by one to hold the input signal.
Sampling speed is determined by the delay buffer interval of each stage.
Trigger
stage

## Slide 10: Evaluation of Analog Memory Cell
Sampling speed：~1Gsps
Dynamic range：３V
Modifications will be done to improve Tdly, Voff
AMC output for an
Input (pulse width~10 nsec)
3V

## Slide 11: Digital filter
Pre-amplifier (ASIC)
Digital Filter (FPGA)
Waveform sampler (AMC)
Detector
(13-inch HAPD)
Noise suppression
f
Flexible
Best effort

## Slide 12: Implementation of Digital filter
Matched optimal filter (MOF) had been implemented in FPGA.
 highly suppress noise (compared to analog filters)
 better S/N, t
signal
noise
signal + noise
Digital
Filter
Better S/N & t
enhanced
suppressed
(Q, T)

## Slide 13: Evaluation of Digital filter
MOF
MOF
Even in a case of S/N ratio = 5 at input,
S/N ratio at output > 10 and Time resolution is < 1nsec

## Slide 14: Goal of FY2006
Evaluation and modification of AMC512 cells
System test (HPD and readout system)