## Slide 1: BURLE INDUSTRIES Recent Photomultiplier and Device Developments
NNN05
Robert Caracciolo
8 April 2005
BURLE INDUSTRIES

## Slide 2: BURLE INDUSTRIES Overview
BURLE INDUSTRIES, INC.
Conversion Tubes
Power Tubes
Real Estate
BURLE ELECTRO-OPTICS, INC.
BURLE INDUSTRIES GmbH
BURLE INDUSTRIES UK LIMITED
BURLE deMexico

## Slide 3: Core Competencies
Conversion Tubes, Lancaster PA
  Conventional PMT design and fabrication
  Photocathode processing
  Image tube design and fabrication
  PMT packaging
  Electronics: VDN, Miniature HVPS, Front-end electronics
Power Tubes, Lancaster PA
  Design and fabrication of vacuum tubes for power generation and switching
  Plating and environmental testing
  Ceramic-to-Metal joining techniques
BEO, Sturbridge MA
  Microchannel plates
  Channel multipliers
  Fiber optics

## Slide 4: PMT Markets
Medical Imaging
  Maintain ~ 30% market share and growing
  Provide high-volume tubes for both SPECT and PET
Have presence in general spectroscopy, scintillation counting, and HEP
Have begun to target the HEP market more aggressively
  Development of the PLANACON family
  Cost competitive fast timing PMTs such as the 8575B.
  SBIR grant to develop large area PMT

## Slide 5: Recent Product Developments
Planacon
Modules
PMT’s

## Slide 6: Planacon™ MCP-PMTs
Two inch square flat PMT with dual MCP multiplier.
Anodes, 2x2 and 8x8 configurations. Additional configurations available.
Bi-alkali cathode on quartz faceplate or cryogenic bi-alkali.
Easily tiled, low profile, photon counting, good time resolution, multi-anode.

## Slide 7: MCP-PMT Construction
Faceplate
Anode & Pins
Indium Seal
Dual MCP
MCP Retainer
Ceramic Insulators
Spacing between faceplate and MCP and MCP and anode can be varied for different applications

## Slide 8: PMT Construction/Processing
Electron multiplier is supported by bulb spacers and leads to the stem
Envelope is evacuated through an exhaust tubulation
Cathode processed in-situ with Sb and alkali dispensers
Tip-off of tubulation using flame or electric heater
Sb bead
Alkali
Channels
Stem
Exhaust tubulation
Bulb
Dynode
Structure
Photocathode

## Slide 9: MCP-PMT Operation
Faceplate
Photocathode
Dual MCP
Anode
Gain ~ 106
Photoelectron
V ~ 200V
V ~ 200V
V ~ 2000V
photon

## Slide 10: Planacon Characteristics
Spatial resolution can be tailored by choice of faceplate, MCP, and pixilated anode
Good photon counting properties at gains of 0.2 – 2 x 106.
Peak to Valley typically > 2:1 with uniform illumination of faceplate.
Output is relatively insensitive to external magnetic fields due to proximity of the cathode, MCP, and anode.
Good pulse height resolution
  10% FWHM, 2” NaI crystal, 662keV.
Cathode uniformity within 10% over full active area.
Anode uniformity ~ 1.5:1 over the 2” active area in analog mode.
Goal is to obtain 1.2:1 anode uniformity.
Cross-talk < 1%.

## Slide 11: Planacon Characteristics (cont’d)
25m pore size with 32 m pitch, investigating the use of 10 m pore size.
40:1 L/D ratio, probably moving to 60:1 for the 10 m devices
Gains of up to 106 with current MCPs
Extended dynamic range glass
Gain is very stable up to ~3% of strip current
Chevron configuration

## Slide 12: Improved Open Area Ratio Planacon
Packaging is streamlined to maximize detector area relative to device dimensions
2.28” sq. vs. 2.50” sq.
0.45” vs. 0.65” ht.
86% vs. 66% OAR
68 gms vs. 128 gms

## Slide 13: Future Directions
Increase the anode configurations offered
Improve Open Area Ratio for tiling applications
Develop variants optimized for
  Photon counting with high spatial resolution
  Low cross-talk and magnetic field immunity
  Cryogenic Applications
  Ultra-low background
Develop other geometries as required by specific markets and applications

## Slide 14: Recent PMT and Module Offerings
8575B, a low cost variant of the 8575.
  Window material is 8250
  Assembly technique is simplified for improved manufacturability
8575Q, Quartz faceplate for UV applications
8575B-800, 8575B module with integrated HVPS and divider
83092 module, short 1” tube for oil well logging applications

## Slide 15: 8575B-800 Module

## Slide 16: 8575B-800 Characteristics
Vacuum potted – ideal for high altitude balloon payloads
Low noise ( 10mV )
Low power (12V @ 1mA)
Regulated
Voltage or resistance controlled

## Slide 17: Large Area PMT Program
Actively working on Phase II objectives of a DOE SBIR to develop a 20” diameter PMT with cost < $0.75/cm2 of active area, including VDN and cabling
Will also develop 2”, 5”, and 8” variants
Want to establish close ties with researchers associated with proton-decay and neutrino experiments to aid in development
Represents a BURLE commitment to becoming a major player in the HEP market

## Slide 18: Requirements
Parameter | Value | Units | Comments
Spectral Response | 300 - 650 | nm | Response < 300nm not very useful due to attenuation length in water
Cathode QE at 390nm | 20 | % | Desire as high as possible
Collection Efficiency | 70 | % | Desire as high as possible
Gain | 1 x 107 |  | 
Dark Counts | 25 | kcps | Desire 3 – 4 kHz at 30C
Transit Time Spread (FWHM) | 5.5 | ns | Desire 3 ns
Photocathode area, head-on | 2000 | cm2 | Sized to give lowest cost per unit area
High Voltage | +2000 | V | Could be higher
Pressure | 8 | atm | Total outside – inside pressure difference. Could use acrylic pressure vessel if needed.
Packaging |  |  | VDN + HV and signal cables, hermetically sealed
Chemical resistance |  |  | Pure H2O

## Slide 19: Photocathode Design
Requirements for highest possible QE and lowest possible dark counts are in conflict.
Trade-study will be performed and initial PMT builds will be designed to optimize these parameters. Dark counts of 3kcps are possible, but QE will probably be limited to 20% max.
Electron multiplier design will influence the dark counts, and will be considered in that design

## Slide 20: Current Activities
Interfacing with glass and bulb manufacturers to optimize cost-effective bulb design and manufacturing approach.
FEA and environmental testing to validate mechanical integrity of bulb.
Employing 2-D and 3-D electron optics models.
  Cathode to Dy1 fields
  Dy1 to the electron multiplier fields
Design and implement novel focusing elements. Required for a bulb with a small neck.
Validated our design concepts on the 2” PMT. Will continue with the 5”, 8”, and 20” PMT’s.
Reviewing different photocathode processes and or design to optimize balance of QE and Dark counts.

## Slide 21: Summary
Established a high volume Manufacturing facility in Mexico to maintain production of PET and SPECT PMT’s
Introducing new product lines to further service both the Medical Imaging Business, as well as other applications including HEP, oil well logging, X-ray digitizers.
Offer Modules to make the application of PMT’s more convenient to the user.
Developing large PMT formats to service the HEP community.