# MEMO

Author: J.E Campagne

PAON2-OptX21cm/11.02.13

Diffusion: Ph. Abbon, R. Ansari, D. Charlet, Ch. Magneville, A. S. Torrentò

Subject: Summary of Electronic Test Bench analysis at LAL (08/02/13)

# Introduction

In references Nançay/Amas/09.02.12 & 13.02.12 and PAON2/04.02.13 problems of oscillations seen in BAOelec data have been presented and the case of thin lines has been addressed on Amas data on references Nançay/Amas/16.03.12 & 04.04.12. It is reminded that in Amas setup the analog & digital electronic chain are coupled to the standard Nançay electronic chain, while for PAON2 this is a stand-alone chain. Besides this difference, the cabling is different, few meters in Amas setup while 50+9m in case of PAON2, and the interconnection too in case of PAON2.

# The setups

To clarify and analyze the different sources of problems it was decided to build a complete test bench with analog and digital chain coupled to the standard BAO DAQ (Figure 1).

Figure 1 Schema of tests analyzed in this MEMO.

Three tests have been analyzed the 8th Feb. 2013: the first one (Test 1) concern the ADC response only, for the second one (Test 2) it is added the mixer-OL while for the last one (Test 3) 7m of coax-cable is added. For all of them, a single channel was used with 8192 x 2 time samples (8192 freq. bins), the trigger was set to 1kHz and RAW data was saved before FFT processing which was done after with spectra averaged every 1sec.

## ADC alone

On Figure 2 is shown the spectra (ie. auto-correlation) over the full frequency band (top panel) and a zoom (bottom panel)

Figure 2 Power in dB of the ADC alone

There are no oscillations modulating the spectra. The 125MHz-line is understood as a FPGA clock which is unavoidable. The other tiny line may also originate from FPGA clock harmonic and/or ADC non linearity artifacts .

## ADC + OL-mixer

On Figure 3 the output of the ADC with OL-mixer (1250MHz) is presented on the full frequency band (top panel) and a zoom (bottom panel).

Figure 3 Power in dB of the ADC + OL-mixer (mind the freq. band)

The spectra don’t show any oscillation pattern, but there are new features concerning thin lines which are potentially a source of limitation of the useful frequency bands:

- A powerful and composite line has emerged around (156.2+1250)MHz

Figure 4

- A series of small line affect the full range [144,153]+1250MHz

Figure 5

- Most of the lines present on Figure 3 have disappeared but some have been enhanced.

## ADC + OL-mixer + 7m cable

The last test differs from the previous one by the addition of 7m coax cable (always terminated by a 50 load). This test is representative of the Amas (and future OptX21cm) setup in the RT chariot. See the results are presented on Figure 6.

Figure 6 Power in dB of the ADC + OL-mixer + 7m cable (mind the freq. band)

The lines present in previous test are also there and the spectra show clearly oscillation distortions with 15MHz p-t-p spacing.

# Summary & Outlook

The test analyzed in this MEMO show clearly the spectra distortions (so-called oscillations) due to the presence of the cable. This may be due to a miss-adaptation of the mixer-OL to the real cable impedance. Also it has been discovered thin lines and sometimes so grouped that an entire frequency band is affected, already present when the OL-mixer is coupled to the ADC.

More investigations are needed to answer questions like: is these phenomena are reproduced for all channels, is there any PLL dis-locking process, etc, and also to study the stability of such effects which may be clearly a limitation of the BAO sensitivity.